## Slide 1: Non-Invasive Ventilator Mask for COVID-19 3D Printing
March 31, 2020

## Slide 2: Requirements
Mask needs to cover all mucus membranes above the neck (eyes, nose, mouth)
Must be sealed and all exhausted air needs to be filtered (N95 or greater)
Needs to be cleanable; Autoclave sterilization is not a requirement
CO2 needs measured in exhaust before filter
Exhaust needs fed back to machine after filter for ventilator design by to atmosphere for CPAP/BiPAP design
Port for administering nebulized meds is desirable
Hot swap of filters

## Slide 3: Scott AV 3000 Series
Primary product used by North American Fire Fighters
Redundant sealing around mouth & nose and around face
Simple universal I/O port with redundant sealing
Comes in 3 different sizes to ensure a good fit; 5 pt head harness
Can be fit with a comms adapter to allow better patient-provider communication
6 week lead time
~$385 (Kevlar) or $290 (Polyester) USD list price

## Slide 4: Scott AV 2000 Series
Primary product used by North American Fire Fighters
Single sealing around mouth, nose and around face
Simple universal I/O port with redundant sealing
Comes in 3 different sizes to ensure a good fit; 4pt head harness
Can be fit with a comms adapter to allow better patient-provider communication
1 week lead time
~$385 (Kevlar) or $290 (Polyester) USD list price

## Slide 5: Adapter interface

## Slide 6: 22 mm Inlet Adapter
Standard interface for ventilator/CPAP/BiPAP equipment
Friction fit of the connecting hose on the tapered male tube interface
Hose Side
Equipment Side

## Slide 7: Functional Interface Diagram - Mask
Scott AV3000 or AV2000 mask
ears
eyes
mouth
nose
From CPAP/BIPAP
22mm ISO 5356-1
Male
>= N95 Filter
Up to 60 l/min
Preferable to find way to combine most into a standard interface with adapters
To CPAP/BIPAP
22mm ISO 5356-1
Male

## Slide 8: Ventilator/CPAP/BiPAP Adapter for AV2000 or AV3000

## Slide 9: Adapter 3D Print Requirements
Printed in ABS or PETG
Initial qty of 100 (London coverage)
Qty up to 1000 (Ontario coverage)

## Slide 10: Functional System Diagram – Ventilator

## Slide 11: Functional System Diagram – CPAP/BiPAP

## Slide 12: System Layout – CPAP/BiPAP

## Slide 13: Adapter R5 - Model
Design incorporates 3D printing DFM analysis, minimized part count, HFE improved latch, and optimization for printing in multiple materials (PETG, ABS)

## Slide 14: Adapter R4 – First ABS Print

## Slide 15: Adapter R4 – First PETG Print

## Slide 16: Adapter 4 – Model Fit-up

## Slide 17: Fit to Mask

## Slide 18: Test Setup Photos
IPAP = 20 cm H2O
EPAP = 5.0 cm H2O
BPM = 12
Ti = 1.0 sec

## Slide 19: Test Pressure Transducer Locations

## Slide 20: Test Results - Overview

## Slide 21: Test Results – BiPAP Ramp Up

## Slide 22: Test Results – Cough Pressure Spike

## Slide 23: Test Results - Breathing Out of Sync

## Slide 24: Test Results - Breathing in Sync

## Slide 25: Test Comments
Large AV3000 mask used
RESMED VPAP III ST-A used
11 days beard growth on test subject
No discernable leakage between mask and face for majority of test
Cough pressure spike did cause a momentary leak between seal and face
There were no leakage alarms during test

## Slide 26: Backup

## Slide 27: Adapter R1 - Model

## Slide 28: Adapter R2 - Model

## Slide 29: Adapter R3
Modelled but not printed or fitted-up.

## Slide 30: Adapter R4

## Slide 31: Confirmed Assumptions
Mask needs to cover all mucus membranes above the neck (eyes, nose, mouth) – Ears not required
The patient will receive nutrition intravenously, if required
All vital signs will be monitored in methods that do not interfere with the function of the mask or its ability to seal against the patients skin
PIP and PEEP functions will be accomplished by the ventilator unit, and not the mask
Autoclave sterilization is not a requirement